## AS-BUILT vs DESIGN
LLAMAS Prototype Red Camera optics -- Design vs. As-Built comparison | [8 empty] | these fields have different units (um instead of fringes for irregularity of ASPheres; fringes of power for radius instead of mm)
these fields are out of the proposed tolerance range
STAT
|  | L1 - fr. ASP | L1 - rea PL | L2 - front | L2 - rear | L3 - front | L3 - rear | L4 - front | L4 - rear | L5 - front | L5 - rear | L6 - fr. PL | L6 - rear | L7 - front | L7 - r. ASP
Radius [mm] | nominal | 135.126 | 6 | 94.207 | 313.978 | 676.644 | 93.378 | 84.215 | 209.488 | 209.488 | 532.166 | 6 | 52.233 | 52.233 | 71.713 | SP rad AVE | SP rad RMS | PL rad AVE | PL rad RMS
| as-built | 135.126 | 1.296 | 94.209 | 314.01 | 676.892 | 93.377 | 84.208 | 209.495 | 209.53 | 532.619 | 1.265 | 52.234 | 52.231 | 71.713 |  |  | 1.2805 | 0.02192 | [fr]
| diff. | 0 | 4.704 | 0.002 | 0.032 | 0.248 | -0.001 | -0.007 | -0.007 | -0.042 | 0.453 | 4.735 | 0.001 | -0.002 | 0 | 0.00994% | 0.03021% |  |  | [%]
Irregularity [fr] | nominal | 1 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 0.75 | 1 | SP irr AVE | SP irr RMS | ASP irr AVE | ASP irr RMS
| as-built | 1.197 | 0.196 | 0.505 | 0.235 | 0.242 | 0.179 | 0.22 | 0.202 | 0.306 | 0.211 | 0.108 | 0.348 | 0.222 | 1.39 | 0.24783 | 0.10074 | 1.2935 | 0.13647 | [fr]
| margin | -0.197 | 0.554 | 0.245 | 0.515 | 0.508 | 0.571 | 0.53 | 0.548 | 0.444 | 0.539 | 0.642 | 0.402 | 0.528 | -0.39
Roughness [A] | nominal | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | SP rough AVE | SP rough RMS | ASP rough AVE | ASP rough RMS
| as-built | 5.917 | 7.048 | 3.976 | 3.433 | 7.672 | 3.693 | 3.965 | 4.005 | 3.658 | 3.734 | 5.904 | 4.563 | 3.814 | 7.252 | 4.6221 | 1.4369 | 6.5845 | 0.94399 | [A]
| margin | 4.083 | 2.952 | 6.024 | 6.567 | 2.328 | 6.307 | 6.035 | 5.995 | 6.342 | 6.266 | 4.096 | 5.437 | 6.186 | 2.748
Central Thick. [mm] | nominal | 10 |  | 4 |  | 22 |  | 25 |  | 5 |  | 6 |  | 24
| as-built | 10.078 |  | 4.046 |  | 22.091 |  | 25.074 |  | 5.079 |  | 6.093 |  | 24.048 |  | AVE | RMS
| diff. | 0.078 |  | 0.046 |  | 0.091 |  | 0.074 |  | 0.079 |  | 0.093 |  | 0.048 |  | 0.072714 | 0.018883 | [mm]
Diameter [mm] | nominal | 70 |  | 70 |  | 94 |  | 89.9 |  | 90 |  | 69.9 |  | 70
| as-built | 70.01 |  | 70.01 |  | 94 |  | 89.923 |  | 90.013 |  | 69.88 |  | 70.02 |  | AVE | RMS
| diff. | 0.01 |  | 0.01 |  | 0 |  | 0.023 |  | 0.013 |  | -0.02 |  | 0.02 |  | 0.008 | 0.014434 | [mm]
EDT [mm] | nominal | 0.01 |  | 0.01 |  | 0.01 |  | 0.01 |  | 0.01 |  | 0.01 |  | 0.01 |  | AVE | RMS
| as-built | 0.004 |  | 0.006 |  | 0.006 |  | 0.001 |  | 0.001 |  | 0.002 |  | 0.003 |  | 0.0032857 | 0.0021381 | [mm]
| margin | 0.006 |  | 0.004 |  | 0.004 |  | 0.009 |  | 0.009 |  | 0.008 |  | 0.007
SAG [mm] if applicable | nominal | n/a | n/a | 5.969 | n/a | 20 | n/a | n/a | n/a | n/a |  | 10.95 | n/a | n/a | 6.41
| as-built | n/a | n/a | 5.978 | n/a | 19.967 | n/a | n/a | n/a | n/a |  | 10.94 | n/a | n/a | 6.445 | AVE | RMS
| diff. |  |  | 0.009 |  | -0.033 |  |  |  |  | 0 | 0.01 |  |  | 0.035 | 0.0042 | 0.02453 | [mm]
OAH [mm] if applicable | nominal | n/a |  | 11.9 |  |  |  | n/a |  |  |  | 16.95 |  | 30.4
| as-built | n/a |  | n/a |  |  |  | n/a |  |  |  | 17.033 |  | 30.458
| diff. |  |  |  |  | 0 |  |  |  | 0 |  | 0.083 |  | 0.058
SFRO [mm] (Surface to Flat RunOut) | nominal | 0.01 | n/a | n/a | 0.02 | n/a | 0.02 | n/1 | n/a | 0.02 | n/a | n/a | 0.1 | n/a | 0.02 | AVE | RMS
| as-built | 0.004 | n/a | n/a | 0.006 | n/a | 0.001 | n/1 | n/a | 0.001 | n/a | n/a | 0.001 | n/a | 0.005 | 0.003 | 0.0022804 | [mm]
| margin | 0.006 |  |  | 0.014 |  | 0.019 |  |  | 0.019 |  |  | 0.099 | #VALUE! | 0.015
PV max/min [fringes] FRONT | max/min | 2.3606 | -1.4236 | 0.3405 | -0.1645 | 0.1473 | -0.0948 | 0.1341 | -0.0854 |  |  | 0.0677 | -0.0406 | 0.0906 | -0.1313
PV max/min [fringes] REAR | max/min | 0.1306 | -0.0656 | 0.084 | 0.1508 | 0.0797 | -0.0995 | 0.0968 | -0.1056 |  |  | 0.2252 | -0.123 | 2.9346 | -1.4588
Mid-Spatial max/min [um] for Asphere, and RMS value [um] | max/min | 0.6279 | -0.5209 | [10 empty] | 2.0809 | -0.7269
| RMS | 0.058 | [11 empty] | 0.072
Coating Run # |  | 2684C | 2688C | 2684C | 2706C | 2684C | 2688C | 2684C |  |  |  | 2680C |  |  | 2684C | SUM | overall AR coating performance (all coated surfaces)
Coating Perfroamnce: Reflection [%] | 560 | 1.332% | 1.574% | 1.332% | 1.434% | 1.332% | 1.574% | 1.332% |  |  |  | 0.9439% |  |  | 1.332% | 88.45%
| 580 | 0.6376% | 0.7838% | 0.6376% | 0.7405% | 0.6376% | 0.7838% | 0.6376% |  |  |  | 0.2323% |  |  | 0.6376% | 94.41%
| 600 | 0.3496% | 0.4042% | 0.3496% | 0.4216% | 0.3496% | 0.4042% | 0.3496% |  |  |  | 0.05054% |  |  | 0.3496% | 97.01%
| 620 | 0.3088% | 0.287% | 0.3088% | 0.3336% | 0.3088% | 0.287% | 0.3088% |  |  |  | 0.133% |  |  | 0.3088% | 97.44%
| 640 | 0.3995% | 0.3068% | 0.3995% | 0.3693% | 0.3995% | 0.3068% | 0.3995% |  |  |  | 0.3106% |  |  | 0.3995% | 96.76%
| 660 | 0.5297% | 0.3905% | 0.5297% | 0.4494% | 0.5297% | 0.3905% | 0.5297% |  |  |  | 0.484% |  |  | 0.5297% | 95.72%
| 680 | 0.6491% | 0.4897% | 0.6491% | 0.5304% | 0.6491% | 0.4897% | 0.6491% |  |  |  | 0.6051% |  |  | 0.6491% | 94.77%
| 700 | 0.7372% | 0.5592% | 0.7372% | 0.5964% | 0.7372% | 0.5592% | 0.7372% |  |  |  | 0.6573% |  |  | 0.7372% | 94.1%
| 720 | 0.7781% | 0.5965% | 0.7781% | 0.6258% | 0.7781% | 0.5965% | 0.7781% |  |  |  | 0.648% |  |  | 0.7781% | 93.82%
| 740 | 0.77% | 0.6014% | 0.77% | 0.6291% | 0.77% | 0.6014% | 0.77% |  |  |  | 0.5882% |  |  | 0.77% | 93.9%
| 760 | 0.7347% | 0.5798% | 0.7347% | 0.5939% | 0.7347% | 0.5798% | 0.7347% |  |  |  | 0.502% |  |  | 0.7347% | 94.22%
| 780 | 0.667% | 0.5382% | 0.667% | 0.5476% | 0.667% | 0.5382% | 0.667% |  |  |  | 0.4024% |  |  | 0.667% | 94.76%
| 800 | 0.5886% | 0.4795% | 0.5886% | 0.4841% | 0.5886% | 0.4795% | 0.5886% |  |  |  | 0.3148% |  |  | 0.5886% | 95.4%
| 820 | 0.4909% | 0.4193% | 0.4909% | 0.4267% | 0.4909% | 0.4193% | 0.4909% |  |  |  | 0.2333% |  |  | 0.4909% | 96.12%
| 840 | 0.4085% | 0.3559% | 0.4085% | 0.3698% | 0.4085% | 0.3559% | 0.4085% |  |  |  | 0.192% |  |  | 0.4085% | 96.73%
| 860 | 0.3477% | 0.3195% | 0.3477% | 0.3262% | 0.3477% | 0.3195% | 0.3477% |  |  |  | 0.1926% |  |  | 0.3477% | 97.14%
| 880 | 0.2789% | 0.2909% | 0.2789% | 0.2962% | 0.2789% | 0.2909% | 0.2789% |  |  |  | 0.2137% |  |  | 0.2789% | 97.54%
| 900 | 0.2647% | 0.3399% | 0.2647% | 0.3102% | 0.2647% | 0.3399% | 0.2647% |  |  |  | 0.2931% |  |  | 0.2647% | 97.42%
| 920 | 0.2392% | 0.3084% | 0.2392% | 0.295% | 0.2392% | 0.3084% | 0.2392% |  |  |  | 0.4268% |  |  | 0.2392% | 97.49%
| 940 | 0.2501% | 0.3261% | 0.2501% | 0.3454% | 0.2501% | 0.3261% | 0.2501% |  |  |  | 0.56% |  |  | 0.2501% | 97.23%
| 960 | 0.2701% | 0.3671% | 0.2701% | 0.3781% | 0.2701% | 0.3671% | 0.2701% |  |  |  | 0.7358% |  |  | 0.2701% | 96.85%
| 980 | 0.3296% | 0.4319% | 0.3296% | 0.4242% | 0.3296% | 0.4319% | 0.3296% |  |  |  | 0.9514% |  |  | 0.3296% | 96.18%
R_avg on 350-570nm [%] |  | 0.4841 | 0.36 | 0.42 | 0.45 | 0.34 |  | 0.34 |  |  | 0.42 | 0.47 |  |  | 0.4841
DOUBLET OAH [mm] | nominal | [7 empty] | 31.505 |  |  |  | 36.4
| as-built | [7 empty] | 31.689 |  |  |  | 36.624
| diff. | [7 empty] | 0.184 |  |  |  | 0.224
DOUBLET SFRO [mm] | nominal | [7 empty] | 0.025 |  |  |  | 0.01
| as-built | [7 empty] | 0.001 |  |  |  | 0.005
| margin | [7 empty] | 0.024 |  |  |  | 0.005
cement thickness | [mm] | [7 empty] | 0.024 |  |  |  | 0.024
| 560 | 0.98668 | 0.98426 | 0.98668 | 0.98566 | 0.98668 | 0.98426 | 0.98668 | 1 | 1 | 1 | 0.99056 | 1 | 1 | 0.98668 | 0.88453
| 580 | 0.99362 | 0.99216 | 0.99362 | 0.99259 | 0.99362 | 0.99216 | 0.99362 | 1 | 1 | 1 | 0.99768 | 1 | 1 | 0.99362 | 0.94414
| 600 | 0.9965 | 0.99596 | 0.9965 | 0.99578 | 0.9965 | 0.99596 | 0.9965 | 1 | 1 | 1 | 0.99949 | 1 | 1 | 0.9965 | 0.97011
| 620 | 0.99691 | 0.99713 | 0.99691 | 0.99666 | 0.99691 | 0.99713 | 0.99691 | 1 | 1 | 1 | 0.99867 | 1 | 1 | 0.99691 | 0.97445
| 640 | 0.99601 | 0.99693 | 0.99601 | 0.99631 | 0.99601 | 0.99693 | 0.99601 | 1 | 1 | 1 | 0.99689 | 1 | 1 | 0.99601 | 0.96757
| 660 | 0.9947 | 0.9961 | 0.9947 | 0.99551 | 0.9947 | 0.9961 | 0.9947 | 1 | 1 | 1 | 0.99516 | 1 | 1 | 0.9947 | 0.95721
| 680 | 0.99351 | 0.9951 | 0.99351 | 0.9947 | 0.99351 | 0.9951 | 0.99351 | 1 | 1 | 1 | 0.99395 | 1 | 1 | 0.99351 | 0.94765
| 700 | 0.99263 | 0.99441 | 0.99263 | 0.99404 | 0.99263 | 0.99441 | 0.99263 | 1 | 1 | 1 | 0.99343 | 1 | 1 | 0.99263 | 0.94102
| 720 | 0.99222 | 0.99403 | 0.99222 | 0.99374 | 0.99222 | 0.99403 | 0.99222 | 1 | 1 | 1 | 0.99352 | 1 | 1 | 0.99222 | 0.93819
| 740 | 0.9923 | 0.99399 | 0.9923 | 0.99371 | 0.9923 | 0.99399 | 0.9923 | 1 | 1 | 1 | 0.99412 | 1 | 1 | 0.9923 | 0.93901
| 760 | 0.99265 | 0.9942 | 0.99265 | 0.99406 | 0.99265 | 0.9942 | 0.99265 | 1 | 1 | 1 | 0.99498 | 1 | 1 | 0.99265 | 0.94225
| 780 | 0.99333 | 0.99462 | 0.99333 | 0.99452 | 0.99333 | 0.99462 | 0.99333 | 1 | 1 | 1 | 0.99598 | 1 | 1 | 0.99333 | 0.94764
| 800 | 0.99411 | 0.99521 | 0.99411 | 0.99516 | 0.99411 | 0.99521 | 0.99411 | 1 | 1 | 1 | 0.99685 | 1 | 1 | 0.99411 | 0.95396
| 820 | 0.99509 | 0.99581 | 0.99509 | 0.99573 | 0.99509 | 0.99581 | 0.99509 | 1 | 1 | 1 | 0.99767 | 1 | 1 | 0.99509 | 0.96115
| 840 | 0.99591 | 0.99644 | 0.99591 | 0.9963 | 0.99591 | 0.99644 | 0.99591 | 1 | 1 | 1 | 0.99808 | 1 | 1 | 0.99591 | 0.96732
| 860 | 0.99652 | 0.9968 | 0.99652 | 0.99674 | 0.99652 | 0.9968 | 0.99652 | 1 | 1 | 1 | 0.99807 | 1 | 1 | 0.99652 | 0.97141
| 880 | 0.99721 | 0.99709 | 0.99721 | 0.99704 | 0.99721 | 0.99709 | 0.99721 | 1 | 1 | 1 | 0.99786 | 1 | 1 | 0.99721 | 0.97541
| 900 | 0.99735 | 0.9966 | 0.99735 | 0.9969 | 0.99735 | 0.9966 | 0.99735 | 1 | 1 | 1 | 0.99707 | 1 | 1 | 0.99735 | 0.97424
| 920 | 0.99761 | 0.99692 | 0.99761 | 0.99705 | 0.99761 | 0.99692 | 0.99761 | 1 | 1 | 1 | 0.99573 | 1 | 1 | 0.99761 | 0.97494
| 940 | 0.9975 | 0.99674 | 0.9975 | 0.99655 | 0.9975 | 0.99674 | 0.9975 | 1 | 1 | 1 | 0.9944 | 1 | 1 | 0.9975 | 0.97226
| 960 | 0.9973 | 0.99633 | 0.9973 | 0.99622 | 0.9973 | 0.99633 | 0.9973 | 1 | 1 | 1 | 0.99264 | 1 | 1 | 0.9973 | 0.96846
| 980 | 0.9967 | 0.99568 | 0.9967 | 0.99576 | 0.9967 | 0.99568 | 0.9967 | 1 | 1 | 1 | 0.99049 | 1 | 1 | 0.9967 | 0.96178
GLASS
effective central thick. |  |  | EFF. CT |  | EFF. CT |  | EFF. CT |  | EFF. CT |  | EFF. CT |  | EFF. CT |  | EFF. CT | SUM | overall internal glass transmission (all 8 lenses)
|  | 10mm |  | 10mm | [5 empty] | 10mm |  | 10mm |  |  |  | #REF!
Glass Internal Transmission | 340 | 0.989 | #VALUE! | 0.81 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.89 | #VALUE! | 0.977 | #VALUE! |  | #VALUE! | #REF!
| 360 | 0.994 | #VALUE! | 0.98 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.987 | #VALUE! | 0.993 | #VALUE! |  | #VALUE! | #REF!
| 380 | 0.998 | #VALUE! | 0.991 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.996 | #VALUE! | 0.997 | #VALUE! |  | #VALUE! | #REF!
| 400 | 0.998 | #VALUE! | 0.997 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.998 | #VALUE! | 0.998 | #VALUE! |  | #VALUE! | #REF!
| 420 | 0.999 | #VALUE! | 0.998 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.998 | #VALUE! | 0.998 | #VALUE! |  | #VALUE! | #REF!
| 440 | 0.999 | #VALUE! | 0.998 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.998 | #VALUE! | 0.998 | #VALUE! |  | #VALUE! | #REF!
| 460 | 0.999 | #VALUE! | 0.998 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.998 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! | #REF!
| 480 | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! | #REF!
| 500 | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! | #REF!
| 520 | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! | #REF!
| 540 | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! | #REF!
| 560 | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! | #REF!
| 580 | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! |  | #VALUE! | 0.999 | #VALUE! | 0.999 | #VALUE! |  | #VALUE! | #REF!

## 2684C
nm | %R
1100 | 1.061 | 1100 | 1.061% | 500 | 7.269%
1099 | 1.0512 | 1080 | 0.897% | 520 | 4.587%
1098 | 1.0391 | 1060 | 0.7567% | 540 | 2.604%
1097 | 1.0309 | 1040 | 0.6227% | 560 | 1.332%
1096 | 1.028 | 1020 | 0.5135% | 580 | 0.6376%
1095 | 1.0144 | 1000 | 0.4074% | 600 | 0.3496%
1094 | 1.0001 | 980 | 0.3296% | 620 | 0.3088%
1093 | 0.99803 | 960 | 0.2701% | 640 | 0.3995%
1092 | 0.99072 | 940 | 0.2501% | 660 | 0.5297%
1091 | 0.99441 | 920 | 0.2392% | 680 | 0.6491%
1090 | 0.97723 | 900 | 0.2647% | 700 | 0.7372%
1089 | 0.96817 | 880 | 0.2789% | 720 | 0.7781%
1088 | 0.96772 | 860 | 0.3477% | 740 | 0.77%
1087 | 0.95537 | 840 | 0.4085% | 760 | 0.7347%
1086 | 0.94275 | 820 | 0.4909% | 780 | 0.667%
1085 | 0.93701 | 800 | 0.5886% | 800 | 0.5886%
1084 | 0.92932 | 780 | 0.667% | 820 | 0.4909%
1083 | 0.92441 | 760 | 0.7347% | 840 | 0.4085%
1082 | 0.91008 | 740 | 0.77% | 860 | 0.3477%
1081 | 0.91342 | 720 | 0.7781% | 880 | 0.2789%
1080 | 0.89701 | 700 | 0.7372% | 900 | 0.2647%
1079 | 0.89599 | 680 | 0.6491% | 920 | 0.2392%
1078 | 0.87839 | 660 | 0.5297% | 940 | 0.2501%
1077 | 0.86821 | 640 | 0.3995% | 960 | 0.2701%
1076 | 0.87008 | 620 | 0.3088% | 980 | 0.3296%
1075 | 0.86282 | 600 | 0.3496% | 1000 | 0.4074%
1074 | 0.85444 | 580 | 0.6376% | 1020 | 0.5135%
1073 | 0.84776 | 560 | 1.332% | 1040 | 0.6227%
1072 | 0.83755 | 540 | 2.604% | 1060 | 0.7567%
1071 | 0.83245 | 520 | 4.587% | 1080 | 0.897%
1070 | 0.82631 | 500 | 7.269% | 1100 | 1.061%
1069 | 0.82097 | 480 | 10.65% | #REF! | #REF!
1068 | 0.81376 | 460 | 14.14% | #REF! | #REF!
1067 | 0.80232 | 440 | 17.34% | #REF! | #REF!
1066 | 0.79634 | 420 | 19.47% | #REF! | #REF!
1065 | 0.78977 | 400 | 19.81% | #REF! | #REF!
1064 | 0.7767 | 0 | 0% | #REF! | #REF!
1063 | 0.78064 | 0 | 0% | #REF! | #REF!
1062 | 0.76753 | 0 | 0% | #REF! | #REF!
1061 | 0.75992 | 0 | 0% | #REF! | #REF!
1060 | 0.75666
1059 | 0.75048
1058 | 0.73541
1057 | 0.73282
1056 | 0.72691
1055 | 0.72333
1054 | 0.70574
1053 | 0.70419
1052 | 0.7028
1051 | 0.6968
1050 | 0.68141
1049 | 0.67725
1048 | 0.67028
1047 | 0.6645
1046 | 0.66134
1045 | 0.6579
1044 | 0.64353
1043 | 0.6428
1042 | 0.62954
1041 | 0.6302
1040 | 0.62273
1039 | 0.61109
1038 | 0.60375
1037 | 0.60256
1036 | 0.5931
1035 | 0.58819
1034 | 0.57995
1033 | 0.57491
1032 | 0.5706
1031 | 0.56408
1030 | 0.5559
1029 | 0.56001
1028 | 0.55279
1027 | 0.54239
1026 | 0.53871
1025 | 0.53377
1024 | 0.52745
1023 | 0.525
1022 | 0.51381
1021 | 0.51424
1020 | 0.51355
1019 | 0.50322
1018 | 0.49136
1017 | 0.48618
1016 | 0.48684
1015 | 0.48047
1014 | 0.46688
1013 | 0.47582
1012 | 0.46447
1011 | 0.45766
1010 | 0.44867
1009 | 0.45069
1008 | 0.45075
1007 | 0.43738
1006 | 0.42962
1005 | 0.4316
1004 | 0.43621
1003 | 0.41984
1002 | 0.41122
1001 | 0.41657
1000 | 0.40736
999 | 0.40695
998 | 0.39174
997 | 0.39972
996 | 0.39658
995 | 0.39177
994 | 0.37716
993 | 0.37918
992 | 0.3755
991 | 0.37413
990 | 0.36896
989 | 0.37394
988 | 0.36873
987 | 0.34903
986 | 0.34354
985 | 0.35643
984 | 0.33913
983 | 0.34161
982 | 0.33703
981 | 0.32636
980 | 0.32958
979 | 0.32942
978 | 0.31614
977 | 0.31405
976 | 0.31942
975 | 0.31028
974 | 0.31763
973 | 0.29871
972 | 0.31416
971 | 0.30421
970 | 0.28125
969 | 0.29049
968 | 0.30624
967 | 0.29618
966 | 0.27844
965 | 0.28963
964 | 0.28808
963 | 0.29421
962 | 0.27438
961 | 0.27958
960 | 0.27009
959 | 0.26683
958 | 0.27621
957 | 0.26813
956 | 0.26617
955 | 0.26045
954 | 0.26454
953 | 0.26821
952 | 0.2643
951 | 0.25254
950 | 0.25035
949 | 0.25363
948 | 0.26694
947 | 0.25752
946 | 0.24771
945 | 0.25556
944 | 0.24941
943 | 0.22588
942 | 0.23403
941 | 0.23915
940 | 0.25008
939 | 0.2491
938 | 0.25438
937 | 0.23804
936 | 0.24939
935 | 0.24006
934 | 0.24266
933 | 0.24756
932 | 0.24035
931 | 0.23105
930 | 0.21183
929 | 0.23682
928 | 0.26125
927 | 0.22384
926 | 0.23099
925 | 0.25437
924 | 0.25953
923 | 0.25579
922 | 0.22924
921 | 0.23233
920 | 0.23916
919 | 0.20521
918 | 0.21925
917 | 0.23015
916 | 0.25837
915 | 0.207
914 | 0.22033
913 | 0.22197
912 | 0.27958
911 | 0.24624
910 | 0.23698
909 | 0.3051
908 | 0.27775
907 | 0.23427
906 | 0.25792
905 | 0.27227
904 | 0.29111
903 | 0.21453
902 | 0.2636
901 | 0.25137
900 | 0.26466
899 | 0.27762
898 | 0.23585
897 | 0.27864
896 | 0.25705
895 | 0.24888
894 | 0.26184
893 | 0.24199
892 | 0.23725
891 | 0.21408
890 | 0.23972
889 | 0.26393
888 | 0.26106
887 | 0.26706
886 | 0.26871
885 | 0.26797
884 | 0.27449
883 | 0.27767
882 | 0.27009
881 | 0.28662
880 | 0.27893
879 | 0.28348
878 | 0.28074
877 | 0.28934
876 | 0.29756
875 | 0.29391
874 | 0.30098
873 | 0.29757
872 | 0.30163
871 | 0.30233
870 | 0.3132
869 | 0.30941
868 | 0.30991
867 | 0.32176
866 | 0.31802
865 | 0.32835
864 | 0.32345
863 | 0.33493
862 | 0.34071
861 | 0.33745
860 | 0.34768
859 | 0.34903
858 | 0.3448
857 | 0.35416
856 | 0.35107
855 | 0.36209
854 | 0.36071
853 | 0.3606
852 | 0.37017
851 | 0.37145
850 | 0.37396
849 | 0.37507
848 | 0.37403
847 | 0.38416
846 | 0.38993
845 | 0.39644
844 | 0.39749
843 | 0.40002
842 | 0.40028
841 | 0.40638
840 | 0.40851
839 | 0.40808
838 | 0.41928
837 | 0.42395
836 | 0.44279
835 | 0.43349
834 | 0.436
833 | 0.44147
832 | 0.44342
831 | 0.45227
830 | 0.45433
829 | 0.46803
828 | 0.4573
827 | 0.46198
826 | 0.47748
825 | 0.47286
824 | 0.47852
823 | 0.4838
822 | 0.4933
821 | 0.49214
820 | 0.49091
819 | 0.49864
818 | 0.51156
817 | 0.50965
816 | 0.50861
815 | 0.52475
814 | 0.52802
813 | 0.53734
812 | 0.53431
811 | 0.53841
810 | 0.54297
809 | 0.54157
808 | 0.54469
807 | 0.54958
806 | 0.55503
805 | 0.56407
804 | 0.56756
803 | 0.57529
802 | 0.57691
801 | 0.58472
800 | 0.58858
799 | 0.59268
798 | 0.60041
797 | 0.59954
796 | 0.60649
795 | 0.60691
794 | 0.60669
793 | 0.61513
792 | 0.61739
791 | 0.62202
790 | 0.62895
789 | 0.63589
788 | 0.6336
787 | 0.63502
786 | 0.64083
785 | 0.64388
784 | 0.65099
783 | 0.65667
782 | 0.66531
781 | 0.66464
780 | 0.66703
779 | 0.66956
778 | 0.67263
777 | 0.67249
776 | 0.67688
775 | 0.68684
774 | 0.69013
773 | 0.69471
772 | 0.69815
771 | 0.69929
770 | 0.69958
769 | 0.70312
768 | 0.70635
767 | 0.71154
766 | 0.71567
765 | 0.71467
764 | 0.72047
763 | 0.72454
762 | 0.72951
761 | 0.73256
760 | 0.73465
759 | 0.7362
758 | 0.74039
757 | 0.74379
756 | 0.7447
755 | 0.74959
754 | 0.74913
753 | 0.7522
752 | 0.75611
751 | 0.7553
750 | 0.75942
749 | 0.75901
748 | 0.75908
747 | 0.75988
746 | 0.76617
745 | 0.77195
744 | 0.77354
743 | 0.77331
742 | 0.77112
741 | 0.77084
740 | 0.77001
739 | 0.77328
738 | 0.77606
737 | 0.77769
736 | 0.77677
735 | 0.77748
734 | 0.77789
733 | 0.78049
732 | 0.78074
731 | 0.7762
730 | 0.78038
729 | 0.77962
728 | 0.7826
727 | 0.78273
726 | 0.77742
725 | 0.78128
724 | 0.78014
723 | 0.78149
722 | 0.77876
721 | 0.77596
720 | 0.77806
719 | 0.77679
718 | 0.77556
717 | 0.77269
716 | 0.77222
715 | 0.77187
714 | 0.76977
713 | 0.77009
712 | 0.76667
711 | 0.76337
710 | 0.76374
709 | 0.76038
708 | 0.75747
707 | 0.75578
706 | 0.75403
705 | 0.75018
704 | 0.75005
703 | 0.74637
702 | 0.74361
701 | 0.74067
700 | 0.73721
699 | 0.73271
698 | 0.73128
697 | 0.7269
696 | 0.72482
695 | 0.7202
694 | 0.71267
693 | 0.71087
692 | 0.70739
691 | 0.7009
690 | 0.69796
689 | 0.69352
688 | 0.68698
687 | 0.68283
686 | 0.68375
685 | 0.67941
684 | 0.66878
683 | 0.66269
682 | 0.65795
681 | 0.6548
680 | 0.64911
679 | 0.64654
678 | 0.63897
677 | 0.63255
676 | 0.6265
675 | 0.61581
674 | 0.61393
673 | 0.61234
672 | 0.60352
671 | 0.59912
670 | 0.59614
669 | 0.58836
668 | 0.57872
667 | 0.57558
666 | 0.5684
665 | 0.56161
664 | 0.55431
663 | 0.54886
662 | 0.54085
661 | 0.53412
660 | 0.52965
659 | 0.52127
658 | 0.5179
657 | 0.50793
656 | 0.49985
655 | 0.49398
654 | 0.48811
653 | 0.48616
652 | 0.47537
651 | 0.46948
650 | 0.46106
649 | 0.45394
648 | 0.45229
647 | 0.44187
646 | 0.43668
645 | 0.4304
644 | 0.42555
643 | 0.41689
642 | 0.41046
641 | 0.4041
640 | 0.39946
639 | 0.39347
638 | 0.38827
637 | 0.38189
636 | 0.37555
635 | 0.37052
634 | 0.36395
633 | 0.35882
632 | 0.35393
631 | 0.35049
630 | 0.34524
629 | 0.3421
628 | 0.33677
627 | 0.33153
626 | 0.32692
625 | 0.32484
624 | 0.32346
623 | 0.31865
622 | 0.31487
621 | 0.31367
620 | 0.30877
619 | 0.30637
618 | 0.30528
617 | 0.30667
616 | 0.30582
615 | 0.30662
614 | 0.30421
613 | 0.30384
612 | 0.30748
611 | 0.30617
610 | 0.30543
609 | 0.31059
608 | 0.31179
607 | 0.31441
606 | 0.3153
605 | 0.3236
604 | 0.32625
603 | 0.33213
602 | 0.33717
601 | 0.34093
600 | 0.34962
599 | 0.35673
598 | 0.3645
597 | 0.37308
596 | 0.37961
595 | 0.39301
594 | 0.40517
593 | 0.41404
592 | 0.4244
591 | 0.43742
590 | 0.45178
589 | 0.46775
588 | 0.48256
587 | 0.49776
586 | 0.517
585 | 0.53295
584 | 0.55324
583 | 0.56985
582 | 0.5931
581 | 0.6151
580 | 0.63758
579 | 0.66025
578 | 0.68475
577 | 0.71034
576 | 0.73721
575 | 0.76641
574 | 0.79521
573 | 0.8251
572 | 0.85659
571 | 0.8899
570 | 0.92262
569 | 0.95758
568 | 0.99532
567 | 1.0322
566 | 1.0714
565 | 1.1124
564 | 1.1522
563 | 1.1955
562 | 1.2383
561 | 1.2871
560 | 1.3319
559 | 1.3812
558 | 1.4322
557 | 1.4836
556 | 1.5349
555 | 1.589
554 | 1.6453
553 | 1.7057
552 | 1.7656
551 | 1.8256
550 | 1.8869
549 | 1.9514
548 | 2.0158
547 | 2.0834
546 | 2.1546
545 | 2.225
544 | 2.2973
543 | 2.3717
542 | 2.4474
541 | 2.5255
540 | 2.6045
539 | 2.6848
538 | 2.7691
537 | 2.8565
536 | 2.9437
535 | 3.0319
534 | 3.122
533 | 3.2158
532 | 3.3099
531 | 3.4046
530 | 3.5017
529 | 3.6013
528 | 3.7023
527 | 3.8059
526 | 3.913
525 | 4.0214
524 | 4.131
523 | 4.2428
522 | 4.357
521 | 4.4712
520 | 4.5865
519 | 4.7032
518 | 4.8223
517 | 4.9446
516 | 5.0674
515 | 5.1931
514 | 5.3202
513 | 5.4479
512 | 5.5765
511 | 5.7072
510 | 5.8403
509 | 5.9756
508 | 6.1113
507 | 6.25
506 | 6.3904
505 | 6.5295
504 | 6.6719
503 | 6.8213
502 | 6.9726
501 | 7.1171
500 | 7.2687
499 | 7.4232
498 | 7.5769
497 | 7.7369
496 | 7.9002
495 | 8.0598
494 | 8.2224
493 | 8.3917
492 | 8.5527
491 | 8.7189
490 | 8.8917
489 | 9.0617
488 | 9.2325
487 | 9.4107
486 | 9.5878
485 | 9.7633
484 | 9.9446
483 | 10.1225
482 | 10.3023
481 | 10.48
480 | 10.6537
479 | 10.8213
478 | 10.9899
477 | 11.1639
476 | 11.335
475 | 11.5041
474 | 11.6896
473 | 11.8648
472 | 12.0433
471 | 12.2187
470 | 12.3903
469 | 12.5636
468 | 12.7435
467 | 12.9188
466 | 13.093
465 | 13.2624
464 | 13.442
463 | 13.6182
462 | 13.7887
461 | 13.9698
460 | 14.1389
459 | 14.3082
458 | 14.4873
457 | 14.6628
456 | 14.8132
455 | 14.981
454 | 15.159
453 | 15.3247
452 | 15.5009
451 | 15.6651
450 | 15.8296
449 | 15.9856
448 | 16.1472
447 | 16.2953
446 | 16.4657
445 | 16.6184
444 | 16.7621
443 | 16.9189
442 | 17.0663
441 | 17.2009
440 | 17.3429
439 | 17.4841
438 | 17.6129
437 | 17.748
436 | 17.8749
435 | 17.9971
434 | 18.1207
433 | 18.2382
432 | 18.342
431 | 18.4726
430 | 18.5641
429 | 18.6719
428 | 18.7919
427 | 18.8854
426 | 18.9934
425 | 19.0756
424 | 19.1581
423 | 19.2425
422 | 19.3456
421 | 19.4047
420 | 19.467
419 | 19.5242
418 | 19.5878
417 | 19.6399
416 | 19.6988
415 | 19.7343
414 | 19.7882
413 | 19.8275
412 | 19.8518
411 | 19.8805
410 | 19.8908
409 | 19.921
408 | 19.938
407 | 19.9365
406 | 19.9371
405 | 19.9185
404 | 19.9251
403 | 19.8962
402 | 19.8754
401 | 19.832
400 | 19.8124

## 2688C
nm | %R
1100 | 1.1693 | 1100 | 1.169% | 500 | 7.709%
1099 | 1.1609 | 1080 | 1.019% | 520 | 4.983%
1098 | 1.1593 | 1060 | 0.8729% | 540 | 2.934%
1097 | 1.1484 | 1040 | 0.7366% | 560 | 1.574%
1096 | 1.1444 | 1020 | 0.6188% | 580 | 0.7838%
1095 | 1.1321 | 1000 | 0.5112% | 600 | 0.4042%
1094 | 1.1225 | 980 | 0.4319% | 620 | 0.287%
1093 | 1.1209 | 960 | 0.3671% | 640 | 0.3068%
1092 | 1.1175 | 940 | 0.3261% | 660 | 0.3905%
1091 | 1.1001 | 920 | 0.3084% | 680 | 0.4897%
1090 | 1.0941 | 900 | 0.3399% | 700 | 0.5592%
1089 | 1.0861 | 880 | 0.2909% | 720 | 0.5965%
1088 | 1.0782 | 860 | 0.3195% | 740 | 0.6014%
1087 | 1.0767 | 840 | 0.3559% | 760 | 0.5798%
1086 | 1.0629 | 820 | 0.4193% | 780 | 0.5382%
1085 | 1.0487 | 800 | 0.4795% | 800 | 0.4795%
1084 | 1.0469 | 780 | 0.5382% | 820 | 0.4193%
1083 | 1.042 | 760 | 0.5798% | 840 | 0.3559%
1082 | 1.041 | 740 | 0.6014% | 860 | 0.3195%
1081 | 1.0284 | 720 | 0.5965% | 880 | 0.2909%
1080 | 1.0192 | 700 | 0.5592% | 900 | 0.3399%
1079 | 1.0039 | 680 | 0.4897% | 920 | 0.3084%
1078 | 0.99942 | 660 | 0.3905% | 940 | 0.3261%
1077 | 1.0036 | 640 | 0.3068% | 960 | 0.3671%
1076 | 0.98379 | 620 | 0.287% | 980 | 0.4319%
1075 | 0.98006 | 600 | 0.4042% | 1000 | 0.5112%
1074 | 0.96986 | 580 | 0.7838% | 1020 | 0.6188%
1073 | 0.96097 | 560 | 1.574% | 1040 | 0.7366%
1072 | 0.9619 | 540 | 2.934% | 1060 | 0.8729%
1071 | 0.94812 | 520 | 4.983% | 1080 | 1.019%
1070 | 0.94182 | 500 | 7.709% | 1100 | 1.169%
1069 | 0.93524 | 480 | 11.1% | #REF! | #REF!
1068 | 0.9262 | 460 | 14.6% | #REF! | #REF!
1067 | 0.92556 | 440 | 17.82% | #REF! | #REF!
1066 | 0.91246 | 420 | 20.01% | #REF! | #REF!
1065 | 0.90707 | 400 | 20.5% | #REF! | #REF!
1064 | 0.89782 | 380 | 18.61% | #REF! | #REF!
1063 | 0.88615 | 360 | 14.16% | #REF! | #REF!
1062 | 0.88929 | 340 | 7.86% | #REF! | #REF!
1061 | 0.88594 | 320 | 2.23% | #REF! | #REF!
1060 | 0.87287
1059 | 0.86576
1058 | 0.86341
1057 | 0.85042
1056 | 0.84352
1055 | 0.82481
1054 | 0.82011
1053 | 0.81792
1052 | 0.80976
1051 | 0.80407
1050 | 0.80375
1049 | 0.79128
1048 | 0.78383
1047 | 0.78062
1046 | 0.77707
1045 | 0.76861
1044 | 0.76171
1043 | 0.761
1042 | 0.75836
1041 | 0.75245
1040 | 0.73657
1039 | 0.73282
1038 | 0.72609
1037 | 0.71736
1036 | 0.71868
1035 | 0.70859
1034 | 0.70208
1033 | 0.68965
1032 | 0.69179
1031 | 0.68326
1030 | 0.67146
1029 | 0.66879
1028 | 0.66865
1027 | 0.66809
1026 | 0.65588
1025 | 0.6438
1024 | 0.64689
1023 | 0.63638
1022 | 0.63418
1021 | 0.62699
1020 | 0.61882
1019 | 0.61082
1018 | 0.60194
1017 | 0.59431
1016 | 0.59055
1015 | 0.59595
1014 | 0.5914
1013 | 0.58814
1012 | 0.58007
1011 | 0.5654
1010 | 0.56484
1009 | 0.55862
1008 | 0.54745
1007 | 0.54658
1006 | 0.54593
1005 | 0.53799
1004 | 0.53248
1003 | 0.52413
1002 | 0.52517
1001 | 0.51496
1000 | 0.5112
999 | 0.51401
998 | 0.50494
997 | 0.50052
996 | 0.50258
995 | 0.49774
994 | 0.4904
993 | 0.48717
992 | 0.48458
991 | 0.48176
990 | 0.47267
989 | 0.46003
988 | 0.46089
987 | 0.45866
986 | 0.45384
985 | 0.44615
984 | 0.43878
983 | 0.44149
982 | 0.44367
981 | 0.44192
980 | 0.43192
979 | 0.4212
978 | 0.41941
977 | 0.41858
976 | 0.40879
975 | 0.41109
974 | 0.41168
973 | 0.4007
972 | 0.39184
971 | 0.38206
970 | 0.379
969 | 0.38327
968 | 0.38093
967 | 0.38061
966 | 0.38812
965 | 0.38579
964 | 0.3847
963 | 0.37063
962 | 0.3685
961 | 0.37353
960 | 0.36705
959 | 0.3726
958 | 0.3629
957 | 0.35282
956 | 0.36297
955 | 0.35733
954 | 0.34517
953 | 0.34479
952 | 0.34571
951 | 0.34481
950 | 0.34996
949 | 0.3321
948 | 0.33667
947 | 0.34061
946 | 0.33508
945 | 0.32938
944 | 0.3237
943 | 0.32894
942 | 0.3201
941 | 0.30856
940 | 0.32609
939 | 0.33113
938 | 0.31438
937 | 0.31794
936 | 0.32859
935 | 0.3138
934 | 0.31124
933 | 0.31273
932 | 0.30766
931 | 0.31492
930 | 0.29171
929 | 0.28559
928 | 0.29984
927 | 0.31169
926 | 0.30891
925 | 0.32292
924 | 0.2907
923 | 0.28531
922 | 0.26776
921 | 0.29278
920 | 0.30837
919 | 0.32644
918 | 0.30917
917 | 0.30389
916 | 0.28293
915 | 0.28395
914 | 0.31348
913 | 0.34088
912 | 0.30305
911 | 0.28404
910 | 0.2829
909 | 0.2876
908 | 0.31393
907 | 0.30867
906 | 0.34386
905 | 0.30937
904 | 0.30388
903 | 0.29496
902 | 0.27807
901 | 0.29257
900 | 0.33986
899 | 0.32233
898 | 0.29985
897 | 0.34003
896 | 0.3767
895 | 0.31662
894 | 0.25918
893 | 0.34576
892 | 0.2919
891 | 0.28409
890 | 0.31062
889 | 0.28292
888 | 0.27846
887 | 0.28082
886 | 0.27981
885 | 0.28451
884 | 0.27883
883 | 0.28272
882 | 0.28997
881 | 0.2871
880 | 0.29087
879 | 0.28962
878 | 0.29882
877 | 0.29293
876 | 0.28694
875 | 0.29007
874 | 0.2946
873 | 0.29772
872 | 0.29335
871 | 0.29751
870 | 0.29502
869 | 0.30656
868 | 0.30764
867 | 0.30685
866 | 0.29573
865 | 0.30348
864 | 0.3159
863 | 0.31405
862 | 0.31548
861 | 0.30918
860 | 0.3195
859 | 0.32732
858 | 0.32475
857 | 0.32432
856 | 0.32805
855 | 0.3314
854 | 0.3375
853 | 0.3377
852 | 0.34256
851 | 0.34607
850 | 0.33911
849 | 0.3458
848 | 0.34829
847 | 0.34109
846 | 0.34574
845 | 0.34807
844 | 0.35412
843 | 0.36252
842 | 0.36918
841 | 0.36437
840 | 0.35588
839 | 0.36759
838 | 0.3772
837 | 0.3838
836 | 0.37716
835 | 0.37789
834 | 0.38189
833 | 0.38021
832 | 0.38716
831 | 0.39025
830 | 0.39428
829 | 0.40053
828 | 0.40258
827 | 0.40049
826 | 0.40816
825 | 0.40555
824 | 0.41185
823 | 0.41268
822 | 0.41554
821 | 0.42136
820 | 0.41934
819 | 0.42326
818 | 0.43136
817 | 0.43821
816 | 0.4396
815 | 0.43175
814 | 0.44572
813 | 0.44361
812 | 0.44851
811 | 0.45157
810 | 0.4588
809 | 0.45792
808 | 0.4586
807 | 0.45739
806 | 0.46035
805 | 0.45846
804 | 0.472
803 | 0.47429
802 | 0.48065
801 | 0.47658
800 | 0.47945
799 | 0.48009
798 | 0.48918
797 | 0.49877
796 | 0.49202
795 | 0.50036
794 | 0.49677
793 | 0.50049
792 | 0.50757
791 | 0.50666
790 | 0.50784
789 | 0.50537
788 | 0.51305
787 | 0.52024
786 | 0.52127
785 | 0.52406
784 | 0.5296
783 | 0.52865
782 | 0.53546
781 | 0.53841
780 | 0.53817
779 | 0.54671
778 | 0.54638
777 | 0.54457
776 | 0.54309
775 | 0.54876
774 | 0.55093
773 | 0.55445
772 | 0.5581
771 | 0.55281
770 | 0.56173
769 | 0.56558
768 | 0.56556
767 | 0.56787
766 | 0.56872
765 | 0.57118
764 | 0.57604
763 | 0.58087
762 | 0.57685
761 | 0.57629
760 | 0.57984
759 | 0.58402
758 | 0.58383
757 | 0.585
756 | 0.58924
755 | 0.58885
754 | 0.58939
753 | 0.59176
752 | 0.5963
751 | 0.59683
750 | 0.59667
749 | 0.59713
748 | 0.59681
747 | 0.59761
746 | 0.59702
745 | 0.60455
744 | 0.602
743 | 0.59962
742 | 0.60354
741 | 0.60186
740 | 0.60142
739 | 0.60122
738 | 0.60524
737 | 0.60505
736 | 0.60628
735 | 0.61162
734 | 0.60793
733 | 0.60915
732 | 0.6099
731 | 0.60699
730 | 0.60314
729 | 0.60315
728 | 0.60508
727 | 0.60771
726 | 0.60589
725 | 0.60014
724 | 0.60326
723 | 0.60401
722 | 0.60318
721 | 0.59885
720 | 0.59653
719 | 0.59645
718 | 0.59643
717 | 0.59405
716 | 0.59246
715 | 0.59467
714 | 0.59309
713 | 0.58976
712 | 0.59189
711 | 0.58794
710 | 0.58274
709 | 0.5802
708 | 0.57992
707 | 0.57618
706 | 0.57562
705 | 0.57476
704 | 0.57039
703 | 0.56677
702 | 0.56263
701 | 0.56459
700 | 0.55924
699 | 0.56044
698 | 0.55495
697 | 0.55251
696 | 0.54902
695 | 0.54068
694 | 0.54225
693 | 0.5375
692 | 0.53616
691 | 0.53267
690 | 0.52757
689 | 0.52573
688 | 0.51743
687 | 0.51792
686 | 0.51557
685 | 0.50435
684 | 0.5031
683 | 0.50149
682 | 0.49771
681 | 0.48999
680 | 0.48973
679 | 0.48786
678 | 0.48032
677 | 0.47696
676 | 0.4675
675 | 0.46626
674 | 0.45662
673 | 0.45605
672 | 0.45288
671 | 0.44921
670 | 0.44192
669 | 0.4366
668 | 0.43134
667 | 0.42846
666 | 0.42578
665 | 0.42007
664 | 0.41478
663 | 0.41161
662 | 0.40421
661 | 0.40006
660 | 0.39049
659 | 0.392
658 | 0.38621
657 | 0.38278
656 | 0.37793
655 | 0.37043
654 | 0.36842
653 | 0.35832
652 | 0.36154
651 | 0.35172
650 | 0.34597
649 | 0.34482
648 | 0.34293
647 | 0.33614
646 | 0.32939
645 | 0.32513
644 | 0.32419
643 | 0.31956
642 | 0.31735
641 | 0.3141
640 | 0.30676
639 | 0.30402
638 | 0.30708
637 | 0.29956
636 | 0.29836
635 | 0.29484
634 | 0.29768
633 | 0.29041
632 | 0.29031
631 | 0.28635
630 | 0.28225
629 | 0.28531
628 | 0.28425
627 | 0.28142
626 | 0.28199
625 | 0.28478
624 | 0.28469
623 | 0.28254
622 | 0.28406
621 | 0.2838
620 | 0.28704
619 | 0.28768
618 | 0.28618
617 | 0.2927
616 | 0.29635
615 | 0.30117
614 | 0.30246
613 | 0.308
612 | 0.31072
611 | 0.31383
610 | 0.32189
609 | 0.32695
608 | 0.33299
607 | 0.34113
606 | 0.34533
605 | 0.35419
604 | 0.36459
603 | 0.37238
602 | 0.38458
601 | 0.39422
600 | 0.40423
599 | 0.41471
598 | 0.42714
597 | 0.43707
596 | 0.45453
595 | 0.4681
594 | 0.48541
593 | 0.49741
592 | 0.51659
591 | 0.53535
590 | 0.553
589 | 0.57124
588 | 0.58903
587 | 0.61118
586 | 0.6337
585 | 0.65546
584 | 0.6801
583 | 0.70565
582 | 0.73196
581 | 0.75551
580 | 0.78383
579 | 0.81126
578 | 0.84142
577 | 0.87051
576 | 0.90273
575 | 0.93313
574 | 0.96836
573 | 1.0059
572 | 1.0395
571 | 1.0797
570 | 1.1179
569 | 1.1577
568 | 1.1983
567 | 1.2407
566 | 1.2857
565 | 1.3287
564 | 1.3753
563 | 1.4212
562 | 1.4716
561 | 1.5238
560 | 1.574
559 | 1.6289
558 | 1.6819
557 | 1.7364
556 | 1.7962
555 | 1.8555
554 | 1.9158
553 | 1.9782
552 | 2.0435
551 | 2.1081
550 | 2.1742
549 | 2.2416
548 | 2.3108
547 | 2.3849
546 | 2.4587
545 | 2.5328
544 | 2.6081
543 | 2.6884
542 | 2.768
541 | 2.8512
540 | 2.934
539 | 3.0194
538 | 3.1077
537 | 3.1966
536 | 3.286
535 | 3.3785
534 | 3.4742
533 | 3.57
532 | 3.6674
531 | 3.7672
530 | 3.8696
529 | 3.9723
528 | 4.0765
527 | 4.1851
526 | 4.2926
525 | 4.4037
524 | 4.5174
523 | 4.6309
522 | 4.7458
521 | 4.8632
520 | 4.9828
519 | 5.1032
518 | 5.2251
517 | 5.3477
516 | 5.4754
515 | 5.6025
514 | 5.7311
513 | 5.8639
512 | 5.9951
511 | 6.1267
510 | 6.2618
509 | 6.3993
508 | 6.5379
507 | 6.6762
506 | 6.8184
505 | 6.9603
504 | 7.1056
503 | 7.2542
502 | 7.406
501 | 7.5583
500 | 7.7085
499 | 7.8663
498 | 8.0248
497 | 8.1848
496 | 8.3456
495 | 8.5078
494 | 8.6667
493 | 8.8347
492 | 9.0038
491 | 9.1708
490 | 9.3464
489 | 9.5189
488 | 9.692
487 | 9.8646
486 | 10.0406
485 | 10.2167
484 | 10.3946
483 | 10.5749
482 | 10.7525
481 | 10.9275
480 | 11.1049
479 | 11.2733
478 | 11.4399
477 | 11.6076
476 | 11.7856
475 | 11.9653
474 | 12.1421
473 | 12.3187
472 | 12.4934
471 | 12.6677
470 | 12.845
469 | 13.0251
468 | 13.2016
467 | 13.3749
466 | 13.5482
465 | 13.724
464 | 13.8997
463 | 14.0727
462 | 14.2476
461 | 14.4293
460 | 14.5979
459 | 14.7659
458 | 14.9408
457 | 15.1108
456 | 15.2873
455 | 15.4525
454 | 15.6175
453 | 15.7913
452 | 15.9521
451 | 16.1312
450 | 16.2885
449 | 16.4496
448 | 16.608
447 | 16.7743
446 | 16.9348
445 | 17.0865
444 | 17.2354
443 | 17.3886
442 | 17.5386
441 | 17.6734
440 | 17.8171
439 | 17.9587
438 | 18.0933
437 | 18.2298
436 | 18.3713
435 | 18.4997
434 | 18.6219
433 | 18.7475
432 | 18.8593
431 | 18.9646
430 | 19.0946
429 | 19.1975
428 | 19.3042
427 | 19.4039
426 | 19.4986
425 | 19.6128
424 | 19.6968
423 | 19.7803
422 | 19.8576
421 | 19.9447
420 | 20.014
419 | 20.0984
418 | 20.1645
417 | 20.2069
416 | 20.2833
415 | 20.3235
414 | 20.3797
413 | 20.4095
412 | 20.4541
411 | 20.4902
410 | 20.5053
409 | 20.5378
408 | 20.5514
407 | 20.5502
406 | 20.5607
405 | 20.555
404 | 20.5715
403 | 20.5603
402 | 20.5429
401 | 20.5232
400 | 20.5034
399 | 20.4526
398 | 20.4159
397 | 20.3755
396 | 20.3248
395 | 20.2706
394 | 20.1958
393 | 20.1311
392 | 20.0391
391 | 19.9763
390 | 19.9024
389 | 19.7835
388 | 19.6828
387 | 19.5732
386 | 19.44
385 | 19.3337
384 | 19.2002
383 | 19.0643
382 | 18.9113
381 | 18.7762
380 | 18.6126
379 | 18.4522
378 | 18.2752
377 | 17.9809
376 | 17.8596
375 | 17.6471
374 | 17.504
373 | 17.3018
372 | 17.103
371 | 16.874
370 | 16.6054
369 | 16.4262
368 | 16.2025
367 | 15.9774
366 | 15.6965
365 | 15.5037
364 | 15.2486
363 | 14.9975
362 | 14.7166
361 | 14.428
360 | 14.1552
359 | 13.8487
358 | 13.5777
357 | 13.267
356 | 12.9809
355 | 12.6591
354 | 12.3637
353 | 12.042
352 | 11.7278
351 | 11.4264
350 | 11.1021
349 | 10.7848
348 | 10.4642
347 | 10.1255
346 | 9.8247
345 | 9.5057
344 | 9.1609
343 | 8.818
342 | 8.5195
341 | 8.2074
340 | 7.8604
339 | 7.5319
338 | 7.2246
337 | 6.8972
336 | 6.5769
335 | 6.2936
334 | 5.9786
333 | 5.6493
332 | 5.3428
331 | 5.0073
330 | 4.7841
329 | 4.4673
328 | 4.2014
327 | 3.9196
326 | 3.6483
325 | 3.3802
324 | 3.1172
323 | 2.9415
322 | 2.7024
321 | 2.4928
320 | 2.2305

## 2706C
nm | %R
1100 | 1.1446 | 1100 | 1.145% | 500 | 6.996%
1099 | 1.1294 | 1080 | 0.9887% | 520 | 4.495%
1098 | 1.1191 | 1060 | 0.8529% | 540 | 2.641%
1097 | 1.1138 | 1040 | 0.7194% | 560 | 1.434%
1096 | 1.1068 | 1020 | 0.6072% | 580 | 0.7405%
1095 | 1.0983 | 1000 | 0.5177% | 600 | 0.4216%
1094 | 1.0866 | 980 | 0.4242% | 620 | 0.3336%
1093 | 1.0807 | 960 | 0.3781% | 640 | 0.3693%
1092 | 1.0825 | 940 | 0.3454% | 660 | 0.4494%
1091 | 1.0724 | 920 | 0.295% | 680 | 0.5304%
1090 | 1.0584 | 900 | 0.3102% | 700 | 0.5964%
1089 | 1.0534 | 880 | 0.2962% | 720 | 0.6258%
1088 | 1.0514 | 860 | 0.3262% | 740 | 0.6291%
1087 | 1.0347 | 840 | 0.3698% | 760 | 0.5939%
1086 | 1.0328 | 820 | 0.4267% | 780 | 0.5476%
1085 | 1.0192 | 800 | 0.4841% | 800 | 0.4841%
1084 | 1.0219 | 780 | 0.5476% | 820 | 0.4267%
1083 | 1.0099 | 760 | 0.5939% | 840 | 0.3698%
1082 | 1.0002 | 740 | 0.6291% | 860 | 0.3262%
1081 | 0.99592 | 720 | 0.6258% | 880 | 0.2962%
1080 | 0.98869 | 700 | 0.5964% | 900 | 0.3102%
1079 | 0.97912 | 680 | 0.5304% | 920 | 0.295%
1078 | 0.97325 | 660 | 0.4494% | 940 | 0.3454%
1077 | 0.96146 | 640 | 0.3693% | 960 | 0.3781%
1076 | 0.96206 | 620 | 0.3336% | 980 | 0.4242%
1075 | 0.95138 | 600 | 0.4216% | 1000 | 0.5177%
1074 | 0.94158 | 580 | 0.7405% | 1020 | 0.6072%
1073 | 0.93927 | 560 | 1.434% | 1040 | 0.7194%
1072 | 0.93299 | 540 | 2.641% | 1060 | 0.8529%
1071 | 0.92275 | 520 | 4.495% | 1080 | 0.9887%
1070 | 0.90901 | 500 | 6.996% | 1100 | 1.145%
1069 | 0.91051 | 480 | 10.16% | #REF! | #REF!
1068 | 0.90424 | 460 | 13.48% | #REF! | #REF!
1067 | 0.88862 | 440 | 16.64% | #REF! | #REF!
1066 | 0.88633 | 420 | 18.87% | #REF! | #REF!
1065 | 0.87904 | 400 | 19.57% | #REF! | #REF!
1064 | 0.87648 | 0 | 0% | #REF! | #REF!
1063 | 0.86649 | 0 | 0% | #REF! | #REF!
1062 | 0.85488 | 0 | 0% | #REF! | #REF!
1061 | 0.85447 | 0 | 0% | #REF! | #REF!
1060 | 0.85291
1059 | 0.83685
1058 | 0.82919
1057 | 0.82791
1056 | 0.82035
1055 | 0.80997
1054 | 0.80548
1053 | 0.79587
1052 | 0.79895
1051 | 0.78885
1050 | 0.77977
1049 | 0.77561
1048 | 0.78645
1047 | 0.76661
1046 | 0.75644
1045 | 0.74638
1044 | 0.7467
1043 | 0.73657
1042 | 0.72403
1041 | 0.72669
1040 | 0.71942
1039 | 0.71289
1038 | 0.70524
1037 | 0.69935
1036 | 0.69842
1035 | 0.69272
1034 | 0.68128
1033 | 0.67199
1032 | 0.68173
1031 | 0.67398
1030 | 0.65387
1029 | 0.65678
1028 | 0.65255
1027 | 0.63846
1026 | 0.63369
1025 | 0.63428
1024 | 0.63045
1023 | 0.6241
1022 | 0.61116
1021 | 0.61274
1020 | 0.60719
1019 | 0.59864
1018 | 0.58667
1017 | 0.58936
1016 | 0.58827
1015 | 0.58278
1014 | 0.57949
1013 | 0.56301
1012 | 0.56129
1011 | 0.56514
1010 | 0.55358
1009 | 0.54481
1008 | 0.5448
1007 | 0.5246
1006 | 0.51674
1005 | 0.5282
1004 | 0.53782
1003 | 0.52901
1002 | 0.5175
1001 | 0.51988
1000 | 0.51769
999 | 0.49835
998 | 0.49602
997 | 0.4881
996 | 0.48527
995 | 0.47316
994 | 0.47403
993 | 0.48038
992 | 0.47341
991 | 0.46657
990 | 0.46591
989 | 0.46991
988 | 0.46464
987 | 0.45
986 | 0.44287
985 | 0.43922
984 | 0.42916
983 | 0.42627
982 | 0.43377
981 | 0.41665
980 | 0.42419
979 | 0.4199
978 | 0.41682
977 | 0.42024
976 | 0.41726
975 | 0.40344
974 | 0.40804
973 | 0.40831
972 | 0.41339
971 | 0.39874
970 | 0.38442
969 | 0.39129
968 | 0.39759
967 | 0.38133
966 | 0.37298
965 | 0.37157
964 | 0.37943
963 | 0.37311
962 | 0.37272
961 | 0.37809
960 | 0.37813
959 | 0.36482
958 | 0.34734
957 | 0.34644
956 | 0.35019
955 | 0.35564
954 | 0.34642
953 | 0.3313
952 | 0.34602
951 | 0.34235
950 | 0.34075
949 | 0.33698
948 | 0.33095
947 | 0.33539
946 | 0.32676
945 | 0.32694
944 | 0.31972
943 | 0.31334
942 | 0.32017
941 | 0.32879
940 | 0.34544
939 | 0.3174
938 | 0.31478
937 | 0.32039
936 | 0.32122
935 | 0.31397
934 | 0.29938
933 | 0.30516
932 | 0.29375
931 | 0.29154
930 | 0.28695
929 | 0.29689
928 | 0.30601
927 | 0.28432
926 | 0.28042
925 | 0.3184
924 | 0.3209
923 | 0.27778
922 | 0.28157
921 | 0.30787
920 | 0.29499
919 | 0.26844
918 | 0.26979
917 | 0.27825
916 | 0.28966
915 | 0.28953
914 | 0.26828
913 | 0.26572
912 | 0.31161
911 | 0.34479
910 | 0.29984
909 | 0.30675
908 | 0.30355
907 | 0.27108
906 | 0.27856
905 | 0.2996
904 | 0.29218
903 | 0.29693
902 | 0.31849
901 | 0.30708
900 | 0.31016
899 | 0.27257
898 | 0.28243
897 | 0.32741
896 | 0.31551
895 | 0.31051
894 | 0.32569
893 | 0.35747
892 | 0.41539
891 | 0.28383
890 | 0.28399
889 | 0.28447
888 | 0.28633
887 | 0.285
886 | 0.2853
885 | 0.28592
884 | 0.28867
883 | 0.29175
882 | 0.29087
881 | 0.28973
880 | 0.29615
879 | 0.29399
878 | 0.28986
877 | 0.28676
876 | 0.29594
875 | 0.29685
874 | 0.30512
873 | 0.29919
872 | 0.29908
871 | 0.29622
870 | 0.31186
869 | 0.31065
868 | 0.30847
867 | 0.31166
866 | 0.30223
865 | 0.32063
864 | 0.31671
863 | 0.31438
862 | 0.32674
861 | 0.31994
860 | 0.32618
859 | 0.32461
858 | 0.33803
857 | 0.33086
856 | 0.33058
855 | 0.32157
854 | 0.33651
853 | 0.33575
852 | 0.3413
851 | 0.34325
850 | 0.34736
849 | 0.34376
848 | 0.35672
847 | 0.34491
846 | 0.35532
845 | 0.36272
844 | 0.36731
843 | 0.36378
842 | 0.3519
841 | 0.36851
840 | 0.36982
839 | 0.36993
838 | 0.3744
837 | 0.38781
836 | 0.38254
835 | 0.38433
834 | 0.39053
833 | 0.38539
832 | 0.38394
831 | 0.3936
830 | 0.38965
829 | 0.39364
828 | 0.40232
827 | 0.4055
826 | 0.40642
825 | 0.41136
824 | 0.41497
823 | 0.42168
822 | 0.41485
821 | 0.42463
820 | 0.42674
819 | 0.43471
818 | 0.43211
817 | 0.43323
816 | 0.43953
815 | 0.44552
814 | 0.44361
813 | 0.43755
812 | 0.45269
811 | 0.46188
810 | 0.4591
809 | 0.46161
808 | 0.46297
807 | 0.45826
806 | 0.46856
805 | 0.475
804 | 0.47759
803 | 0.48091
802 | 0.48042
801 | 0.4779
800 | 0.48411
799 | 0.48994
798 | 0.49269
797 | 0.49466
796 | 0.49815
795 | 0.50331
794 | 0.50463
793 | 0.50658
792 | 0.51697
791 | 0.51895
790 | 0.51895
789 | 0.51952
788 | 0.52588
787 | 0.52598
786 | 0.52389
785 | 0.52865
784 | 0.5348
783 | 0.53951
782 | 0.54458
781 | 0.54818
780 | 0.54759
779 | 0.54462
778 | 0.54617
777 | 0.55645
776 | 0.56069
775 | 0.55825
774 | 0.56405
773 | 0.56509
772 | 0.56653
771 | 0.56932
770 | 0.56703
769 | 0.57228
768 | 0.57622
767 | 0.57677
766 | 0.581
765 | 0.58583
764 | 0.58949
763 | 0.58796
762 | 0.58877
761 | 0.59621
760 | 0.59387
759 | 0.59343
758 | 0.59608
757 | 0.59906
756 | 0.60779
755 | 0.60279
754 | 0.60585
753 | 0.60972
752 | 0.60718
751 | 0.60957
750 | 0.61489
749 | 0.61384
748 | 0.61419
747 | 0.61695
746 | 0.61732
745 | 0.61886
744 | 0.61918
743 | 0.61697
742 | 0.61886
741 | 0.62566
740 | 0.62907
739 | 0.62787
738 | 0.62494
737 | 0.625
736 | 0.62443
735 | 0.62836
734 | 0.62978
733 | 0.62771
732 | 0.62587
731 | 0.62601
730 | 0.62616
729 | 0.62945
728 | 0.62806
727 | 0.62676
726 | 0.62861
725 | 0.62877
724 | 0.62631
723 | 0.62748
722 | 0.62556
721 | 0.62301
720 | 0.62585
719 | 0.62461
718 | 0.6228
717 | 0.62092
716 | 0.61866
715 | 0.61747
714 | 0.6194
713 | 0.61736
712 | 0.61561
711 | 0.61486
710 | 0.61241
709 | 0.61171
708 | 0.60942
707 | 0.61081
706 | 0.60774
705 | 0.60511
704 | 0.60141
703 | 0.60201
702 | 0.60002
701 | 0.59717
700 | 0.59642
699 | 0.593
698 | 0.59031
697 | 0.59109
696 | 0.58631
695 | 0.58228
694 | 0.5774
693 | 0.57715
692 | 0.57631
691 | 0.56982
690 | 0.56799
689 | 0.56309
688 | 0.56156
687 | 0.55803
686 | 0.55575
685 | 0.55245
684 | 0.54798
683 | 0.54723
682 | 0.5412
681 | 0.53369
680 | 0.53035
679 | 0.53025
678 | 0.52847
677 | 0.52099
676 | 0.51394
675 | 0.51414
674 | 0.50871
673 | 0.50538
672 | 0.49915
671 | 0.49912
670 | 0.49244
669 | 0.48782
668 | 0.48116
667 | 0.48135
666 | 0.47317
665 | 0.4733
664 | 0.46761
663 | 0.46142
662 | 0.45716
661 | 0.4526
660 | 0.44939
659 | 0.44449
658 | 0.43967
657 | 0.4351
656 | 0.43315
655 | 0.42598
654 | 0.42191
653 | 0.41952
652 | 0.41224
651 | 0.40867
650 | 0.40654
649 | 0.39807
648 | 0.39572
647 | 0.39319
646 | 0.38731
645 | 0.38511
644 | 0.38137
643 | 0.37304
642 | 0.37249
641 | 0.36748
640 | 0.36932
639 | 0.36648
638 | 0.35935
637 | 0.3518
636 | 0.35346
635 | 0.34936
634 | 0.34865
633 | 0.34523
632 | 0.34359
631 | 0.34022
630 | 0.33952
629 | 0.33843
628 | 0.33609
627 | 0.33686
626 | 0.33282
625 | 0.33425
624 | 0.33308
623 | 0.33114
622 | 0.32953
621 | 0.33087
620 | 0.33363
619 | 0.33367
618 | 0.33417
617 | 0.33524
616 | 0.33869
615 | 0.34181
614 | 0.34115
613 | 0.3435
612 | 0.35038
611 | 0.3515
610 | 0.35696
609 | 0.36193
608 | 0.36598
607 | 0.37103
606 | 0.37594
605 | 0.38159
604 | 0.39061
603 | 0.39446
602 | 0.4045
601 | 0.4105
600 | 0.42164
599 | 0.42982
598 | 0.4416
597 | 0.45492
596 | 0.46419
595 | 0.47815
594 | 0.49047
593 | 0.5027
592 | 0.51543
591 | 0.53079
590 | 0.54533
589 | 0.56045
588 | 0.57701
587 | 0.59695
586 | 0.61504
585 | 0.63386
584 | 0.65264
583 | 0.67306
582 | 0.6947
581 | 0.71646
580 | 0.74052
579 | 0.76738
578 | 0.79184
577 | 0.81642
576 | 0.84589
575 | 0.8733
574 | 0.90299
573 | 0.93392
572 | 0.96783
571 | 0.99655
570 | 1.0272
569 | 1.0663
568 | 1.1016
567 | 1.1382
566 | 1.1749
565 | 1.2133
564 | 1.2553
563 | 1.2976
562 | 1.3401
561 | 1.3877
560 | 1.4335
559 | 1.4799
558 | 1.5271
557 | 1.5769
556 | 1.6292
555 | 1.6805
554 | 1.7333
553 | 1.7885
552 | 1.846
551 | 1.9024
550 | 1.9607
549 | 2.023
548 | 2.0859
547 | 2.1499
546 | 2.2158
545 | 2.2824
544 | 2.3497
543 | 2.4219
542 | 2.495
541 | 2.5665
540 | 2.6412
539 | 2.718
538 | 2.7963
537 | 2.875
536 | 2.9568
535 | 3.0414
534 | 3.1263
533 | 3.213
532 | 3.3021
531 | 3.3916
530 | 3.4824
529 | 3.5751
528 | 3.6719
527 | 3.7689
526 | 3.8676
525 | 3.9675
524 | 4.0704
523 | 4.1746
522 | 4.2786
521 | 4.3856
520 | 4.4946
519 | 4.6048
518 | 4.7172
517 | 4.8287
516 | 4.9449
515 | 5.062
514 | 5.1776
513 | 5.2977
512 | 5.4198
511 | 5.5409
510 | 5.6627
509 | 5.7896
508 | 5.917
507 | 6.0434
506 | 6.1732
505 | 6.3038
504 | 6.4376
503 | 6.5733
502 | 6.7116
501 | 6.8538
500 | 6.9962
499 | 7.1422
498 | 7.2839
497 | 7.4292
496 | 7.5835
495 | 7.7371
494 | 7.8928
493 | 8.0449
492 | 8.1992
491 | 8.3579
490 | 8.5142
489 | 8.6732
488 | 8.8336
487 | 9.0007
486 | 9.1673
485 | 9.3322
484 | 9.5015
483 | 9.6703
482 | 9.835
481 | 9.9993
480 | 10.163
479 | 10.3245
478 | 10.4837
477 | 10.6461
476 | 10.8084
475 | 10.9732
474 | 11.1393
473 | 11.3025
472 | 11.4745
471 | 11.6442
470 | 11.8064
469 | 11.9773
468 | 12.1453
467 | 12.306
466 | 12.4804
465 | 12.6494
464 | 12.8154
463 | 12.9796
462 | 13.1439
461 | 13.3168
460 | 13.4792
459 | 13.6423
458 | 13.8165
457 | 13.9771
456 | 14.1401
455 | 14.3152
454 | 14.4822
453 | 14.6427
452 | 14.8072
451 | 14.9573
450 | 15.1182
449 | 15.2808
448 | 15.4236
447 | 15.5858
446 | 15.7499
445 | 15.9039
444 | 16.0553
443 | 16.2115
442 | 16.3433
441 | 16.4886
440 | 16.6364
439 | 16.7741
438 | 16.9038
437 | 17.031
436 | 17.1647
435 | 17.2914
434 | 17.4349
433 | 17.5502
432 | 17.6587
431 | 17.7894
430 | 17.8931
429 | 18.0219
428 | 18.1414
427 | 18.2217
426 | 18.3193
425 | 18.4339
424 | 18.5205
423 | 18.6197
422 | 18.7035
421 | 18.7869
420 | 18.868
419 | 18.9506
418 | 19.0258
417 | 19.0852
416 | 19.1516
415 | 19.2236
414 | 19.2667
413 | 19.3268
412 | 19.3758
411 | 19.4074
410 | 19.4481
409 | 19.471
408 | 19.525
407 | 19.5503
406 | 19.5475
405 | 19.5619
404 | 19.5875
403 | 19.5939
402 | 19.5858
401 | 19.5941
400 | 19.5694

## 2680C
nm | %R
1100 | 2.5887 | 1100 | 2.589% | 500 | 9.353%
1099 | 2.5777 | 1080 | 2.285% | 520 | 5.271%
1098 | 2.554 | 1060 | 1.996% | 540 | 2.513%
1097 | 2.5417 | 1040 | 1.708% | 560 | 0.9439%
1096 | 2.5326 | 1020 | 1.444% | 580 | 0.2323%
1095 | 2.5108 | 1000 | 1.18% | 600 | 0.05054%
1094 | 2.4914 | 980 | 0.9514% | 620 | 0.133%
1093 | 2.4781 | 960 | 0.7358% | 640 | 0.3106%
1092 | 2.4649 | 940 | 0.56% | 660 | 0.484%
1091 | 2.4572 | 920 | 0.4268% | 680 | 0.6051%
1090 | 2.4378 | 900 | 0.2931% | 700 | 0.6573%
1089 | 2.4221 | 880 | 0.2137% | 720 | 0.648%
1088 | 2.4082 | 860 | 0.1926% | 740 | 0.5882%
1087 | 2.3964 | 840 | 0.192% | 760 | 0.502%
1086 | 2.3745 | 820 | 0.2333% | 780 | 0.4024%
1085 | 2.3558 | 800 | 0.3148% | 800 | 0.3148%
1084 | 2.3469 | 780 | 0.4024% | 820 | 0.2333%
1083 | 2.3334 | 760 | 0.502% | 840 | 0.192%
1082 | 2.3233 | 740 | 0.5882% | 860 | 0.1926%
1081 | 2.3009 | 720 | 0.648% | 880 | 0.2137%
1080 | 2.2853 | 700 | 0.6573% | 900 | 0.2931%
1079 | 2.2733 | 680 | 0.6051% | 920 | 0.4268%
1078 | 2.2573 | 660 | 0.484% | 940 | 0.56%
1077 | 2.2431 | 640 | 0.3106% | 960 | 0.7358%
1076 | 2.2265 | 620 | 0.133% | 980 | 0.9514%
1075 | 2.2117 | 600 | 0.05054% | 1000 | 1.18%
1074 | 2.1932 | 580 | 0.2323% | 1020 | 1.444%
1073 | 2.1824 | 560 | 0.9439% | 1040 | 1.708%
1072 | 2.165 | 540 | 2.513% | 1060 | 1.996%
1071 | 2.1572 | 520 | 5.271% | 1080 | 2.285%
1070 | 2.1397 | 500 | 9.353% | 1100 | 2.589%
1069 | 2.1272 | 480 | 14.89% | #REF! | #REF!
1068 | 2.1151 | 460 | 20.98% | #REF! | #REF!
1067 | 2.0982 | 440 | 27.05% | #REF! | #REF!
1066 | 2.0835 | 420 | 31.87% | #REF! | #REF!
1065 | 2.0657 | 400 | 34.41% | #REF! | #REF!
1064 | 2.0501 | 0 | 0% | #REF! | #REF!
1063 | 2.04 | 0 | 0% | #REF! | #REF!
1062 | 2.021 | 0 | 0% | #REF! | #REF!
1061 | 2.0115 | 0 | 0% | #REF! | #REF!
1060 | 1.9957
1059 | 1.9857
1058 | 1.967
1057 | 1.9558
1056 | 1.938
1055 | 1.9249
1054 | 1.9105
1053 | 1.8939
1052 | 1.8792
1051 | 1.8645
1050 | 1.8501
1049 | 1.8359
1048 | 1.8205
1047 | 1.8083
1046 | 1.7918
1045 | 1.7843
1044 | 1.7685
1043 | 1.7525
1042 | 1.7379
1041 | 1.7234
1040 | 1.7077
1039 | 1.6916
1038 | 1.6805
1037 | 1.661
1036 | 1.6523
1035 | 1.6417
1034 | 1.6291
1033 | 1.6098
1032 | 1.6004
1031 | 1.5791
1030 | 1.5649
1029 | 1.5546
1028 | 1.5482
1027 | 1.534
1026 | 1.5147
1025 | 1.5049
1024 | 1.4907
1023 | 1.4751
1022 | 1.4651
1021 | 1.4566
1020 | 1.4436
1019 | 1.4259
1018 | 1.4151
1017 | 1.3994
1016 | 1.3877
1015 | 1.3769
1014 | 1.3546
1013 | 1.3477
1012 | 1.3353
1011 | 1.3142
1010 | 1.3046
1009 | 1.2976
1008 | 1.2834
1007 | 1.2696
1006 | 1.2572
1005 | 1.2427
1004 | 1.2404
1003 | 1.2211
1002 | 1.2047
1001 | 1.1936
1000 | 1.18
999 | 1.1765
998 | 1.1536
997 | 1.1501
996 | 1.1353
995 | 1.1218
994 | 1.0988
993 | 1.0922
992 | 1.0832
991 | 1.0672
990 | 1.0592
989 | 1.0622
988 | 1.0374
987 | 1.0121
986 | 1.0078
985 | 1.0124
984 | 0.99158
983 | 0.97791
982 | 0.95863
981 | 0.95415
980 | 0.95141
979 | 0.9386
978 | 0.91929
977 | 0.90836
976 | 0.90225
975 | 0.87642
974 | 0.87464
973 | 0.85969
972 | 0.87021
971 | 0.85202
970 | 0.82381
969 | 0.8262
968 | 0.83012
967 | 0.80813
966 | 0.78023
965 | 0.78789
964 | 0.77108
963 | 0.76877
962 | 0.74809
961 | 0.74794
960 | 0.73583
959 | 0.72298
958 | 0.72247
957 | 0.70931
956 | 0.69621
955 | 0.68817
954 | 0.68601
953 | 0.6864
952 | 0.6665
951 | 0.65382
950 | 0.6409
949 | 0.62662
948 | 0.63395
947 | 0.61746
946 | 0.594
945 | 0.60232
944 | 0.58924
943 | 0.56364
942 | 0.58372
941 | 0.56242
940 | 0.55997
939 | 0.56286
938 | 0.54968
937 | 0.53974
936 | 0.53141
935 | 0.51989
934 | 0.51926
933 | 0.51149
932 | 0.50132
931 | 0.49295
930 | 0.45749
929 | 0.4732
928 | 0.47151
927 | 0.44503
926 | 0.44373
925 | 0.45267
924 | 0.44394
923 | 0.4477
922 | 0.3946
921 | 0.42117
920 | 0.42682
919 | 0.38168
918 | 0.37986
917 | 0.39551
916 | 0.39803
915 | 0.38426
914 | 0.37454
913 | 0.35503
912 | 0.37415
911 | 0.34452
910 | 0.32279
909 | 0.35957
908 | 0.35412
907 | 0.35741
906 | 0.32432
905 | 0.35628
904 | 0.33289
903 | 0.29193
902 | 0.31163
901 | 0.31678
900 | 0.29306
899 | 0.31514
898 | 0.27503
897 | 0.30124
896 | 0.27168
895 | 0.29806
894 | 0.28695
893 | 0.26304
892 | 0.27262
891 | 0.20867
890 | 0.21107
889 | 0.24024
888 | 0.23417
887 | 0.23414
886 | 0.22691
885 | 0.22839
884 | 0.22704
883 | 0.22625
882 | 0.21383
881 | 0.21489
880 | 0.21368
879 | 0.20871
878 | 0.20355
877 | 0.20405
876 | 0.20357
875 | 0.19686
874 | 0.19997
873 | 0.19414
872 | 0.192
871 | 0.18977
870 | 0.19258
869 | 0.19005
868 | 0.18963
867 | 0.19238
866 | 0.19052
865 | 0.19112
864 | 0.18919
863 | 0.19
862 | 0.18885
861 | 0.18631
860 | 0.19262
859 | 0.18375
858 | 0.18811
857 | 0.18253
856 | 0.17777
855 | 0.18429
854 | 0.17865
853 | 0.18787
852 | 0.18603
851 | 0.182
850 | 0.18351
849 | 0.18163
848 | 0.17871
847 | 0.19095
846 | 0.19547
845 | 0.19601
844 | 0.19201
843 | 0.18329
842 | 0.18961
841 | 0.19449
840 | 0.19195
839 | 0.18902
838 | 0.19543
837 | 0.20163
836 | 0.20107
835 | 0.19743
834 | 0.19897
833 | 0.20288
832 | 0.20974
831 | 0.21194
830 | 0.21135
829 | 0.21456
828 | 0.21182
827 | 0.2136
826 | 0.2225
825 | 0.22177
824 | 0.22546
823 | 0.24
822 | 0.23526
821 | 0.23471
820 | 0.23326
819 | 0.23916
818 | 0.24953
817 | 0.2425
816 | 0.25097
815 | 0.26009
814 | 0.26388
813 | 0.26578
812 | 0.26108
811 | 0.26666
810 | 0.26863
809 | 0.2719
808 | 0.27997
807 | 0.28417
806 | 0.28372
805 | 0.29265
804 | 0.29837
803 | 0.29852
802 | 0.30112
801 | 0.30916
800 | 0.3148
799 | 0.32099
798 | 0.32592
797 | 0.3239
796 | 0.32915
795 | 0.33152
794 | 0.33157
793 | 0.34219
792 | 0.34435
791 | 0.35001
790 | 0.35448
789 | 0.36462
788 | 0.36729
787 | 0.36578
786 | 0.37348
785 | 0.3801
784 | 0.38584
783 | 0.38823
782 | 0.39558
781 | 0.39907
780 | 0.40239
779 | 0.40787
778 | 0.41107
777 | 0.41308
776 | 0.42048
775 | 0.42818
774 | 0.43226
773 | 0.4386
772 | 0.44565
771 | 0.45214
770 | 0.4553
769 | 0.45716
768 | 0.46356
767 | 0.4726
766 | 0.47554
765 | 0.47849
764 | 0.48371
763 | 0.48764
762 | 0.49115
761 | 0.49595
760 | 0.50204
759 | 0.50616
758 | 0.51076
757 | 0.51577
756 | 0.52025
755 | 0.52635
754 | 0.52643
753 | 0.52987
752 | 0.53664
751 | 0.54144
750 | 0.54828
749 | 0.55124
748 | 0.55448
747 | 0.55788
746 | 0.56629
745 | 0.5749
744 | 0.57736
743 | 0.57889
742 | 0.58069
741 | 0.58286
740 | 0.58823
739 | 0.59229
738 | 0.59408
737 | 0.59586
736 | 0.59891
735 | 0.60391
734 | 0.60905
733 | 0.61426
732 | 0.61909
731 | 0.61933
730 | 0.62294
729 | 0.62502
728 | 0.62983
727 | 0.63187
726 | 0.63281
725 | 0.63592
724 | 0.63847
723 | 0.64192
722 | 0.64541
721 | 0.64323
720 | 0.64799
719 | 0.6474
718 | 0.65062
717 | 0.65228
716 | 0.65439
715 | 0.65363
714 | 0.65559
713 | 0.65695
712 | 0.65679
711 | 0.65781
710 | 0.65837
709 | 0.65779
708 | 0.65797
707 | 0.66294
706 | 0.66107
705 | 0.65673
704 | 0.65983
703 | 0.66019
702 | 0.66004
701 | 0.65968
700 | 0.65728
699 | 0.65409
698 | 0.65498
697 | 0.6547
696 | 0.65132
695 | 0.64836
694 | 0.64995
693 | 0.64482
692 | 0.6465
691 | 0.64275
690 | 0.64163
689 | 0.63899
688 | 0.63212
687 | 0.63174
686 | 0.62759
685 | 0.62377
684 | 0.62026
683 | 0.6143
682 | 0.61268
681 | 0.61075
680 | 0.60513
679 | 0.60107
678 | 0.59569
677 | 0.59074
676 | 0.58296
675 | 0.57708
674 | 0.57537
673 | 0.57283
672 | 0.56348
671 | 0.56062
670 | 0.55308
669 | 0.54814
668 | 0.54093
667 | 0.53771
666 | 0.52854
665 | 0.52193
664 | 0.51244
663 | 0.50504
662 | 0.49783
661 | 0.4914
660 | 0.48396
659 | 0.47533
658 | 0.47215
657 | 0.4602
656 | 0.45305
655 | 0.44674
654 | 0.43757
653 | 0.43114
652 | 0.42039
651 | 0.41219
650 | 0.40125
649 | 0.39117
648 | 0.38898
647 | 0.37662
646 | 0.36614
645 | 0.35903
644 | 0.34967
643 | 0.33973
642 | 0.32829
641 | 0.32071
640 | 0.31064
639 | 0.30166
638 | 0.2927
637 | 0.28634
636 | 0.27432
635 | 0.26531
634 | 0.25519
633 | 0.24764
632 | 0.23753
631 | 0.22955
630 | 0.2201
629 | 0.21035
628 | 0.20058
627 | 0.19128
626 | 0.18247
625 | 0.17453
624 | 0.16919
623 | 0.15684
622 | 0.14946
621 | 0.14049
620 | 0.13301
619 | 0.12638
618 | 0.11692
617 | 0.1129
616 | 0.10435
615 | 0.099882
614 | 0.092955
613 | 0.086755
612 | 0.082237
611 | 0.07402
610 | 0.07036
609 | 0.066361
608 | 0.061544
607 | 0.0578
606 | 0.051925
605 | 0.055162
604 | 0.05377
603 | 0.049564
602 | 0.048741
601 | 0.047567
600 | 0.050537
599 | 0.048603
598 | 0.051211
597 | 0.053981
596 | 0.055737
595 | 0.062177
594 | 0.067766
593 | 0.072807
592 | 0.074837
591 | 0.085321
590 | 0.092404
589 | 0.10068
588 | 0.11171
587 | 0.12394
586 | 0.13562
585 | 0.14776
584 | 0.16209
583 | 0.17771
582 | 0.19528
581 | 0.2136
580 | 0.23234
579 | 0.25217
578 | 0.2757
577 | 0.29633
576 | 0.32149
575 | 0.34942
574 | 0.37605
573 | 0.4045
572 | 0.4344
571 | 0.46823
570 | 0.50114
569 | 0.53713
568 | 0.57577
567 | 0.61427
566 | 0.65456
565 | 0.69865
564 | 0.74318
563 | 0.79042
562 | 0.83881
561 | 0.89185
560 | 0.94387
559 | 0.99871
558 | 1.0539
557 | 1.1141
556 | 1.1765
555 | 1.24
554 | 1.3071
553 | 1.3763
552 | 1.4487
551 | 1.5234
550 | 1.5983
549 | 1.6771
548 | 1.7589
547 | 1.8428
546 | 1.9302
545 | 2.0196
544 | 2.1117
543 | 2.2091
542 | 2.3077
541 | 2.4093
540 | 2.5132
539 | 2.62
538 | 2.7296
537 | 2.8437
536 | 2.9608
535 | 3.0805
534 | 3.2033
533 | 3.3282
532 | 3.4587
531 | 3.5921
530 | 3.7265
529 | 3.864
528 | 4.0084
527 | 4.154
526 | 4.3033
525 | 4.4582
524 | 4.6135
523 | 4.7731
522 | 4.9372
521 | 5.1019
520 | 5.2714
519 | 5.442
518 | 5.6171
517 | 5.7975
516 | 5.9806
515 | 6.1671
514 | 6.3555
513 | 6.5473
512 | 6.7413
511 | 6.9397
510 | 7.1414
509 | 7.3459
508 | 7.5574
507 | 7.7727
506 | 7.9835
505 | 8.1984
504 | 8.4213
503 | 8.6527
502 | 8.8896
501 | 9.1179
500 | 9.3535
499 | 9.5982
498 | 9.8505
497 | 10.1041
496 | 10.3597
495 | 10.6181
494 | 10.8832
493 | 11.1546
492 | 11.4169
491 | 11.6875
490 | 11.9739
489 | 12.2486
488 | 12.5285
487 | 12.8183
486 | 13.1062
485 | 13.4023
484 | 13.7022
483 | 14.0053
482 | 14.3031
481 | 14.596
480 | 14.8912
479 | 15.1759
478 | 15.458
477 | 15.757
476 | 16.0503
475 | 16.3466
474 | 16.6626
473 | 16.963
472 | 17.2662
471 | 17.5723
470 | 17.8791
469 | 18.1816
468 | 18.4906
467 | 18.8048
466 | 19.1121
465 | 19.4073
464 | 19.721
463 | 20.04
462 | 20.351
461 | 20.663
460 | 20.9807
459 | 21.2817
458 | 21.6063
457 | 21.9226
456 | 22.2175
455 | 22.5415
454 | 22.8413
453 | 23.1674
452 | 23.4955
451 | 23.7949
450 | 24.1033
449 | 24.4081
448 | 24.6955
447 | 24.9922
446 | 25.3111
445 | 25.6295
444 | 25.8988
443 | 26.1922
442 | 26.4767
441 | 26.7669
440 | 27.0486
439 | 27.3399
438 | 27.6138
437 | 27.8882
436 | 28.1505
435 | 28.415
434 | 28.6903
433 | 28.9467
432 | 29.1905
431 | 29.4349
430 | 29.6796
429 | 29.9125
428 | 30.154
427 | 30.3681
426 | 30.6406
425 | 30.8585
424 | 31.0798
423 | 31.2542
422 | 31.4971
421 | 31.7139
420 | 31.8696
419 | 32.0824
418 | 32.2516
417 | 32.4286
416 | 32.5988
415 | 32.7602
414 | 32.9479
413 | 33.1012
412 | 33.2485
411 | 33.3671
410 | 33.5006
409 | 33.6223
408 | 33.7739
407 | 33.8711
406 | 33.9529
405 | 34.0427
404 | 34.1254
403 | 34.2181
402 | 34.2361
401 | 34.3166
400 | 34.4053